عطاء رقم 13/ل ر/2022

توريد وتركيب وتشغيل أجهزة مخبرية/ كلية الطب

Orbital shaker incubator

Item (1)                                                                                                        QTY (2)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Thermofisher scientific | Manufacturer | 1 |
| MaxQ™ 4000 Benchtop Orbital Shakers | Model Number | 2 |
| Latest production (please specify starting production year) | Starting Production Year | 4 |
| FOR Insect cell culture • Extraction procedures • General mixing • Staining and destaining • Hybridizations • Solubility studies • Plasmid production • Bacterial suspensions • Washing procedures • Protein expression | Application | 5 |
| - Clear acrylic lid - Safety interlock | Design | 6 |
| Orbital motion | Shaking Motion | 7 |
| 19mm | Orbit | 8 |
| LED or LCD for speed and time | Display | 9 |
| - Stainless steel Platform for large bottles and flasks | Shaking Surface | 10 |
| - Approx. 460mm x 460mm with non-slip rubber material - 6 positions | Flat Platform Dimensions (D x W) | 11 |
| - From 15 rpm to 400 rpm or better - Increments of 1rpm | Shaking Speed | 12 |
| - ±1rpm | Speed Accuracy | 13 |
| - Programmable Continuous operation | Time | 14 |
| - 15°C below ambient to 55°C or better ±0.1°C | Temperature range and accuracy | 15 |
| - ±0.5°C at 37°C or better | Temperature uniformity | 16 |
| - Stainless steel construction Standard flask clamps for 25ml, 50ml, 250ml, 500ml and 1000ml | Flask Clamps | 17 |
| Over temperature protection | Safety Features | 18 |
| - Should be supply with all needed accessories | All Standard Accessories | 20 |
| - 220V/50-60 Hz | Power Supply | 21 |

E gel electrophoresis system

Item (2)                                                                                                        QTY (1)

Specifications and description:

- Separate DNA in as little as 15 minutes or less

- Allows electrophoresis and image capture all in one

- E gels and necessary startup kits

- 1 electrophoresis device

- 1 camera device

- 1 power adapter

- 1 pair of viewing glasses

Microplate photometer system

Item (3)                                                                                                  QTY (1 reader 1 washer)

Reader

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Latest production (please specify starting production year) | Starting Production Year | 1 |
| Automatic reader compatible with strips and 96-well microplates that have flat, U, or V bottom wells | Plate Types | 2 |
| Absorbance | Detection Mode | 3 |
| 340 to 800nm or better | Wavelength Range | 4 |
| From 0.000 to 4.000Abs | Photometric Range | 5 |
| ±1% | Photometric Accuracy | 6 |
| 0.001 Abs | Resolution | 7 |
| 0.1% | Repeatability | 8 |
| ≤2.0% | Linearity | 9 |
| <0.005Abs/10min | Stability | 10 |
| Should be supplied at least with 3 filters 405nm, 450nm, and 620nm or better | Filters | 11 |
| Microprocessor controlled | Control | 12 |
| Tungsten - halogen lamp | Light Source | 13 |
| LCD | Display | 14 |
| Automatic calibration before each reading | Calibration | 15 |
| Not more than 50sec for whole plate | Reading Speed | 16 |
| Should have facility for storage of calibration curves | Storage | 17 |
| External high quality laser color printer | Printer | 18 |
| - USB | Interface | 19 |
| - 220 V / 50 Hz | Power Supply | 20 |

washer

| No | Technical Specifications | Requirement |
| --- | --- | --- |
| 1 | Starting Production Year | Latest production |
| 2 | Application | Microplate washer |
| 3 | Plate Type | Automatic washer compatible with strips and 96-well microplates that have flat, U, or V bottom wells |
| 4 | Control | Microprocessor controlled |
| 5 | Display | LCD |
| 6 | Programming Modes | - Washing time - Soaking time - Shaking time - Dispensing volume - Liquid flow rate - Vacuum strength - Plate carriage |
| 7 | Wash Program | Up to 50 programmable washing sequences |
| 8 | Wash Cycles | Adjustable |
| 9 | Wash Head | Programmable needle positions (horizontal or vertical) to an accuracy of 0.1 mm for bottom washing, crosswise aspiration, and overflow washing |
| 10 | Manifold | Easily removable 8 way manifolds |
| 11 | Manifold Types | 96-well washing |
| 12 | Residual Aspiration Volume | <2.0 µL/ well |
| 13 | Alarms | - Over flow - Waste level |
| 14 | Vacuum Pump | Integrated |
| 15 | Bottles | Included for waste, wash and rinse |
| 16 | Power Supply | 220 V / 50 Hz |

Balance

Item (4)                                                                                                        QTY (1)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer and country of origin | Manufacturer | 1 |
| Please specify model number of the offered equipment | Model Number | 2 |
| Latest production (please specify starting production year) | Starting Production Year | 3 |
| - Large glass draught shield with 3 sliding - Doors height-adjustable feet - Chemical resistant housing - Chamber eliminates the influence of air flow | Design | 4 |
| LCD for results, errors, time and functions in progress | Display | 5 |
| 200g approx. | Capacity | 6 |
| ±0.1mg or better | Readability | 7 |
| ±0.2mg or better | Linearity | 8 |
| Built-in auto calibration | Calibration | 9 |
| Made of stainless steel | Weighing Pan | 10 |
| Four different units of measure: g, lb, mg, ct | Weighing Units | 11 |
| - Balance ready - Balance level | Indicators | 12 |
| ≤ 4sec. | Stabilization Time | 13 |
| - Overload protection - Auto shut-off of inactivity | Safety | 14 |
| 220V / 50Hz or AC/DC adapter | Power Supply | 15 |

PH Meter

Item (5)									Qty (1)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer and country of origin | Manufacturer | 1 |
| Please specify model number of the offered equipment | Model Number | 2 |
| Latest production (please specify starting production year) | Starting Production Year | 3 |
| - Compact design - Bench top - Microprocessor control - Waterproof housing and chemical resistant construction | Design | 4 |
| - Range 0 – 14 or better - Resolution 0.01 - Accuracy ±0.01 | PH | 5 |
| - Range ±1999 or better - Resolution 0.1 - Accuracy ±0.2 | mV | 6 |
| - Range 0 – 100 or better - Resolution 0.1 - Accuracy ±0.3 | Temperature | 7 |
| Automatic | Temperature Compensation | 8 |
| LCD | Display | 9 |
| Up to 5 points calibration | Calibration | 10 |
| Automatic | Buffer Recognition | 11 |
| Storage of up to 200 results or better | Data Storage | 12 |
| With stand and holder | PH electrode | 13 |
| Low battery | Indicator | 14 |
| Built-in | Rechargeable Battery | 15 |
| 220V / 50Hz or AC/DC adapter | Power Supply | 16 |

Shaking water bath

Item (6)                                                                                                        QTY (1)

| Chamber Material | 304 Stainless Steel |
| --- | --- |
| Controller Type | Digital |
| Description | 27L Shaking Water Bath |
| Display | Monochrome LCD |
| Heating Capacity | 1500 w |
| Includes | Stainless Steel Gable Cover and Rubber Duck |
| Material | Epoxy Powder-coated Cold Rolled Steel |

| Temperature Presets | 4 |
| --- | --- |
| Temperature Stability | ±0.1°C |
| Temperature Uniformity | ±0.05°C |
| Capacity (Metric) | 27 L |
| Electrical Requirements | Voltage V 200/230 V 50/60 Hz |
| Temperature Range | Ambient +5° to 100°C |

Bench CENTRIFUGE (TUBE)

Item (7)                                                                                                        QTY (1)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer ,country of origin and country of manufacturer | Manufacturer | 1 |
| Please specify model number | Model Number | 2 |
| FDA approval or CE marking | Safety Standard | 3 |
| Latest production (please specify starting production year) | Starting Production Year | 4 |
| - Bench top - Heavy duty - Microprocessor control - Brushless motor (inverter preferred) - Motorized (powered) lid locking device - Tachometer view port - Automatic rotor verification(recognition) - Stainless steel rotor chamber | Design | 5 |
| Not less than 10 programs | Programs | 6 |
| LCD for speed, time and program number | Display | 7 |
| - Digital - Duration from 10s or better to 99min - Timer set steps 10s or better - Short time operation mode-hold | Timer | 8 |
| - Speed - Lid open - Start / stop - Quick (short) run | Visual Indicators | 9 |
| - Sound at end of operation - Rotor error | Audible Signal | 10 |
| ≤ 60 db at maximum speed | Noise Level | 11 |
| - 4 place swing out (Please specify rotor structure) - Adjustable speed from 500rpm or better to 4000rpm - Speed set steps 100rpm or better - Max acceleration / deceleration time 40s or better - 16 tubes per rotor - Set of 4 bucket - 4-places in insert (carrier, adapter) - 10ml tube capacity - Blood collection tube type (vacutainer) - Tube dimensions 16 diameter x (100 – 110) length | Rotor | 12 |
| - Emergency lid release - Imbalance switch-off - Lid locking and holding during rotor run - Motor over heating protection - Rotor over speed protection | Safety | 13 |
| 220 V / 50 Hz | Power Supply | 14 |

Microcentauer (microcentrifuge)

Item (8)                                                                                                        QTY (1)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Latest production (please specify starting production year) | Starting Production Year | 1 |
| - Bench top - Heavy duty - Microprocessor control - Brushless motor - Tachometer view port - Automatic lid opening at the end of the run | Design | 2 |
| LCD or LED for time and speed | Display | 3 |
| - Digital - Duration from 30s or better to 99min - Short time operation mode-hold | Timer | 4 |
| - Rpm/rcf - Lid open - Start / stop - Quick (short) run | Visual Indicators | 5 |
| ≤ 55db at max speed | Noise level | 6 |
| - Fixed-angle PCR rotor - Adjustable speed from 1000 or better to 14500 in rpm - Speed set steps 100 in rpm or better - Acceleration time to max speed 30s or better - Deceleration time from max speed 30s or better - 24 tubes per rotor - PCR tubes capacity from 0.2ml to 1.5ml | Rotor | 7 |
| - Emergency lid release - Imbalance switch-off - Lid locking and holding during rotor run - Optimal ventilation flow | Safety | 8 |
| Included with all needed accessories to run the equipment as recommended by manufacturer (adapters, rotor lid, rotor nut..) | Accessories | 9 |
| 220 V / 50 Hz | Power supply | 10 |

Megafuge REFRIGERATED CENTRIFUGE

Item (9)                                                                                                        QTY (1)

| No | Technical Specifications | Requirement |
| --- | --- | --- |
| 1 | Starting Production Year | Latest production (please specify starting production year) |
| 2 | Design | - Bench top - Heavy duty - Microprocessor control - Brushless converter motor - Automatic motorized door lock - Tachometer view port - Automatic rotor verification (recognition) - Stainless steel rotor chamber - Fast cool function (preferable) |
| 3 | Programs | Not less than 20 programs |
| 4 | Display | LCD display for time, speed and temperature |
| 5 | Timer | - Digital - Duration from 1 min or better to 99 min - Timer set steps 1 min or better |
| 6 | Refrigeration System | CFC free |
| 7 | Temperature Range | From -10°C to +40°C |
| 8 | Temperature Accuracy | ±1ºC or better |
| 9 | Alert Indicators | Power failure, imbalance, lid open and over temperature |
| 10 | Rotors | Multiple types of rotors with suitable accessories for every type |
| 11 | Safety | - Emergency lid release - Imbalance switch-off - Lid locking and holding during rotor run |
| 12 | Noise Level | ≤ 60db at max speed |
| 13 | Interface | - USB - LAN |
| 14 | All Standard Accessories | Should be supply with all needed accessories |
| 15 | Power Supply | 220 V/ 50 Hz |

THERMAL CYCLER

Item (10)                                                                                            			QTY (1)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer and country of origin | Manufacturer | 1 |
| Please specify model number of the offered equipment | Model Number | 2 |
| FDA approval or CE marking | Safety standard | 3 |
| Latest production (please specify starting production year) | Starting Production Year | 4 |
| Flexible for all PCR, allows running multiple independent assays | Application | 5 |
| 96 well | Block Capacity | 6 |
| 5.0°C/s | Max. Heating/Cooling rate | 7 |
| 10-30 µl | Sample volume range | 8 |
| 4 - 99C | Temperature range | 9 |
| ±0.25C | Temperature accuracy | 10 |
| Digital | Display | 11 |
| User- friendly | Software | 12 |
| Heating lid indicator | Lid indicator | 13 |
| <0.5 °C | Temperature uniformity | 14 |
| - USB and On-board | Memory | 15 |
| 220V / 50Hz | Power supply | 16 |

LAB REFRIGERATOR

Item (11)                                                                                                        QTY (1)

| No | Technical Specifications | Requirement |
| --- | --- | --- |
| 1 | Design and Construction | - Vertical (upright) - All metal parts inside should be stainless steel - Metal parts from outside made of bacteria-resistant powder coated steel sheet with high quality insulation material between the walls - Hermetically sealed compressor - System should have a positive forced air circulation to maintain temperature uniformity at all levels - Internal LED lighting - Evaporating unit with fan - Condensing unit with fan - Defrost system System should be mounted on four castors with minimum two brakes |
| 2 | Capacity | - 500 liter Approx. |
| 3 | Shelves | - Not less than five stainless steel shelves |
| 4 | Cooling Gas | CFC Free |
| 5 | Display | LCD display for temperature |
| 6 | Temperature Range | From +2ºC to +8ºC or better |
| 7 | Temperature Control | Microprocessor control |
| 8 | Temperature Stability | ± 0.3ºC or better |
| 9 | Temperature Accuracy | ±0.3ºC or better |
| 10 | Audible / Visual Alarms | - Door open - High / low temperature Power failure |
| 11 | Safety | - Alarm system should be incorporated with battery backup Safety thermostat prevents dropping of the temperature below 1°C |
| 12 | Alert Indicators | - Power on - Over temperature |
| 13 | Noise Level | ≤ 60db |
| 14 | Power Supply | - 220 V / 50 Hz |

LAB freezer

Item (12)                                                                                                        QTY (1)

| No | Technical Specifications | Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 3 | Design and Construction | - Vertical (upright) - All metal parts inside should be stainless steel - System should have a positive forced air circulation to maintain temperature uniformity at all levels - Automatic fast defrosting - Internal LED lighting |
| 4 | Capacity | - 230 liter minimum. |
| 5 | Shelves | Not less than five stainless steel shelves |
| 6 | Display | - LCD display for temperature, date and time |
| 7 | Cooling Gas | - CFC free |
| 8 | Temperature Control | Microprocessor controlled |
| 9 | Temperature Range | From within –18ºc to -30ºc |
| 10 | Temperature Stability | ± 1.0ºc or better |
| 11 | Temperature Accuracy | ±0.5ºc or better |
| 12 | Noise Level | ≤ 60db |
| 13 | Condenser Dust Filter | Included |
| 14 | Audible / Visual Alarms | - Door open - High / low temperature Power failure |
| 15 | Alert Indicators | - Power on Over temperature |
| 16 | Power Supply | - 220 V / 50 Hz |

Multichannel pipettes

Item (13)                                                                            QTY (2 different volumes, see table)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer and country of origin | Manufacturer | 1 |
| Please specify model number | Model Number | 2 |
| FDA approval or CE marking | Safety standard | 3 |
| 8 | Number of channels | 4 |
| (30-300) (5-50) µL | Adjustable volume | 5 |
| ± 3% to ± 1% or better | Accuracy | 6 |
| Autoclavable | Manifold | 7 |

Single pipettes adjustable automatic

Item (14)                                                                      QTY (5 SEE VOLUMES in table)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer and country of origin | Manufacturer | 1 |
| Please specify model number | Model Number | 2 |
| Adjustable volume | Type | 3 |
| Digital LCD | Display | 4 |
| (0.2- 2 ) µL, ( 2 - 20 ) µL, ( 20 - 200 ) µL, (200-1000) µL, 200-1000) µL | Volume | 5 |
| Available by the providing company | Maintenance, calibration & parts | 6 |
| ± 1% | Accuracy | 7 |

Single pipette fixed volume

Item (15)                                                                              QTY (5 SEE VOLUMES in table)

| Requirement | Technical Specifications | No |
| --- | --- | --- |
| Please specify manufacturer and country of origin | Manufacturer | 1 |
| Please specify model number of the offered equipment | Model Number | 2 |
| FDA approval or CE marking | Safety standard | 3 |
| (10) µL , (50) µL , (100) µL , (500) µL (1000) µL | Volume | 4 |
| ± 1 % or better | Accuracy | 5 |

Vaccum pump (Pipette pump)

Item (16)                                                                                                        QTY (2)

| No | Technical Specifications | Requirement |
| --- | --- | --- |
| 1 | Manufacturer | Please specify manufacturer and country of origin |
| 2 | Model Number | Please specify model number of the offered equipment |
| 4 | Starting Production Year | Latest production (please specify starting production year) |
| 5 | Technical specifications | - Suitable for all pipettes up to 100 ml - Efficient battery for a long runtime, low battery alert - Fine adjustment of pipetting speed - Tight fit of any pipette into the pipette socket - Handy and light weighted for fatigue-free pipetting - Simple but powerful motor for fast and reliable pipetting - Gravity drain mode for draining out without using any motor - Nosepiece with replaceable filter and silicone adapter - Including lithium -ion battery. - Includes 0.45µm standard hydrophobic membrane filter or better. |
| 6 | Included with the offer | - Calibration to be done following delivery upon request. - To be supplied with compatible filters free of charge |

Counting champer

Item (17)                                                                                                        QTY (2)

Specs:

Depth of chamber: 0.1 mm